MIQE checklist

| Item to check | Importance | Response |
| --- | --- | --- |
| Experimental Design |  |  |
| Definition of experimental and control groups | E | Our control group to which the data has been compared is the V5-eGFP group. The V5-Ago2 and V5-Ago2mut groups are the experimental groups |
| Number within each group | E | 3 technical replicates per biological replicate. 3 biological replicates. |
| Nucleic acid extraction |  |  |
| Nucleic acid extraction procedure | E | RNA extraction from cell culture using TRIzol reagent. |
| Name of kit | E | TRIzol reagent following manufacturer’s protocol (Thermo Fisher Scientific). |
| Details of DNase/RNase treatment | E | No nuclease digestion performed. |
| Contamination assessment | E | Quality assessed on nanodrop. |
| Nucleic acid quantification | E | UV spectrophotometer. |
| Instrument and method | E | Nanodrop RNA measurement. |
| RNA integrity: method/instrument | E | A260/280 measurement. |
| RIN/RQI or Cq of 3’ and 5’ transcripts | E | Not measured, nanodrop used for assessment. |
| Inhibition testing (Cq dilutions, spike or other) | E | Quality was assessed on nanodrop. |
| Reverse Transcription |  |  |
| Complete reaction conditions | E | 1µg of RNA 1µl 10mM dNTPs 1µl Oligo(dT)15 Primer 4µl 5X First strand synthesis buffer 1µl 0.1M DTT 1µl RNase Inhibitor 1µl SuperscriptIII Up to 20µl Nuclease-free water |
| Amount of RNA and reaction volume | E | 1µg of RNA in 20ul reaction volume |
| Priming oligonucleotide and concentration | E | Oligo(dT)15 500µg/ml |
| Reverse transcriptase and concentration | E | Superscript III 200U/µl |
| Temperature and time | E | Denature RNA at 65°C for 5 minutes Reverse transcription at 50°C for 60 minutes Inactivation at 70°C for 15 minutes |
| qPCR target information |  |  |
| Gene symbol | E | S7: RpS7 SFV: SFV4 nsp3 |
| Sequence accession number | E | S7: Vectorbase AAEL009496 SFV4: Genbank KP699763 |
| Amplicon length | E | S7: 202nt SFV:205nt |
| In silico specificity (BLAST etc) | E | BLAST |
| Location of each primer by exon or intron (if applicable) | E | S7: Forward primer in Exon1 and Reverse Primer in Exon 2 No splicing for SFV |
| What splice variants are targeted | E | No alternative splicing in SFV. S7: AAEL009496-RA |
| qPCR oligonucleotides |  |  |
| Primer sequences | E | qS7 F: CCAGGCTATCCTGGAGTTG qS7 R: GACGTGCTTGCCGGAGAAC qSFV F: GCAAGAGGCAAACGAACAGA qSFV R: GGGAAAAGATGAGCAAACCA |
| Location and identity of any modifications | E | No modifications |
| qPCR protocol |  |  |
| Complete Reaction conditions | E | 12.5µl Fast SYBR green master mix 1µl 10µM Forward primer 1µl 10µM Reverse primer 2µl cDNA 8.5µl Nuclease free water |
| Reaction volume and amount of cDNA/DNA | E | 20µl final reaction volume with 2µl cDNA |
| Primer, Mg2+ and dNTP concentrations | E | 400nM primer concentrations Salt and dNTP concentrations are determined by the Fast SYBR green mastermix. |
| Polymerase identity and concentration | E | AmpliTaq Fast DNA polymerase |
| Buffer/kit identity and manufacturer | E | 2X FAST SYBR green reagent (Thermo Fisher Scientific) |
| Additives | E | SYBR green 1 Dye ROX passive reference dye |
| Complete thermocycling parameters | E | Activation step: 95°C for 20 seconds 40 cycles of: Denaturation step: 95°C for 3 seconds Annealing/Extension step: 60°C for 30 seconds |
| Manufacturer of qPCR instrument | E | Applied Biosystems (Thermo Fisher Scientific) |
| qPCR validation |  |  |
| Specificity (gel, sequence, melt or digest) | E | Single melt curve peak. |
| For SYBR green I, Cq of the NTC | E | Undetermined. |
| Calibration curves with slope and y intercept | E | Primers used were previously published. S7: McFarlane et al 2014 SFV: Breakwell et al 2007 |
| PCR efficiency calculated from slope | E | Primers used were previously published. S7: McFarlane et al 2014 SFV: Breakwell et al 2007 |
| R2 of calibration curve | E | Primers used were previously published. S7: McFarlane et al 2014 SFV: Breakwell et al 2007 |
| Linear dynamic range | E | Primers used were previously published. S7: McFarlane et al 2014 SFV: Breakwell et al 2007 |
| Cq variation at LOD | E | Primers used were previously published. S7: McFarlane et al 2014 SFV: Breakwell et al 2007 |
| Evidence of LOD | E | Primers used were previously published. S7: McFarlane et al 2014 SFV: Breakwell et al 2007 |
| If multiplex, efficiency and LOD of each assay | E | N/A |
| Data analysis |  |  |
| qPCR analysis program (source, version) | E | QuantStudio Design & Analysis Software v1.5.3 (Thermo Fisher Scientific) |
| Method of Cq determination | E | Automatic fixed threshold. |
| Outlier identification and disposition | E | No outliers identified via visual inspection of the amplification curve. |
| Results for NTCs | E | NTCs were not included in the analysis as they were undetermined. |
| Justification of number and choice of reference genes | E | Standard commonly used reference gene which is unaffected by virus infection. |
| Description of normalization method | E | Normalisation performed using 2-ΔΔCT method described by Livak and Schmittgen 2001. |
| Number and stage (reverse transcription or qPCR) of technical replicates | E | qPCR technical replicates performed in triplicate. |
| Repeatability (intraassay variation) | E | Variation between assays can be viewed as error bars. |
| Statistical methods for results significance | E | ANOVA |
| Software (source, version) | E | GraphPad Prism v10.2.2 |